## Slide 1
A
B
C
D
Sample
Hi
Lo
Hi
Lo
Hi
Lo
Hi
Lo
1
2
3
4
3
4
Hi
Lo
Hi
Lo
A
B
C
D
A
C
B
D
I+
I-
V+
V-
17
18
25
26
19
20
27
28
(24)
(08)
(41)
(25)
(26)
(10)
(43)
(27)
(35)
(19)
(20)
(04)
#
- Row/Column
#
- Channel
(#)
- Wiring Pin
Wiring Color Code
A – Red
B – Black
C – Blue
D – Green
I – White
V – Yellow
Resistivity/Hall Effect Matrix Card – Keithley 7709
Hi
Lo
Hi
Lo
5
6
A
D
B
C
21
22
29
30
(28)
(12)
(45)
(29)

## Slide 2
Resistivity Measurement Sequence – RAB, CD
Current Off
Short Card – Close 17, 25, 26, 27, 28, 29, 30
Prepare Measurement - Open 25, 27, 28, 29, 30
Current On, Delay, Measure Voltage
Current Reverse, Delay, Measure Voltage
Current Off
Short Card – Close 25, 27, 28, 29, 30

## Slide 3
Resistivity Measurement Sequence – RCD, AB
Current Off
Short Card – Close 17, 25, 26, 27, 28, 29, 30
Prepare Measurement – Close 18
Prepare Measurement – Open 17, 26, 27, 28, 29, 30
Current On, Delay, Measure Voltage
Current Reverse, Delay, Measure Voltage
Current Off
Short Card – Close 17, 26, 27, 28, 29, 30
Finish – Open 18

## Slide 4
Resistivity Measurement Sequence – RAC,BD
Current Off
Short Card – Close 17, 25, 26, 27, 28, 29, 30
Prepare Measurement – Close 19
Prepare Measurement – Open 17, 25, 26, 27, 29, 30
Current On, Delay, Measure Voltage
Current Reverse, Delay, Measure Voltage
Current Off
Short Card – Close 17, 25, 26, 27, 29, 30
Finish – Open 19

## Slide 5
Resistivity Measurement Sequence – RBD,AC
Current Off
Short Card – Close 17, 25, 26, 27, 28, 29, 30
Prepare Measurement – Close 20
Prepare Measurement – Open 17, 25, 26, 28, 29, 30
Current On, Delay, Measure Voltage
Current Reverse, Delay, Measure Voltage
Current Off
Short Card – Close 17, 25, 26, 28, 29, 30
Finish – Open 20

## Slide 6
Hall Effect Measurement Sequence – R±AD,BC
Current Off, B-Field Off
Short Card – Close 17, 25, 26, 27, 28, 29, 30
Prepare Measurement – Close 21
Prepare Measurement – Open 17, 25, 26, 27, 28, 29
Current On, B-Field On, Delay, Measure Voltage
Current Reverse, Delay, Measure Voltage
B-Field Reverse, Delay, Measure Voltage,
Current Reverse, Delay, Measure Voltage
Current Off, B-Field Off
Short Card – Close 17, 25, 26, 27, 28, 29
Finish – Open 21

## Slide 7
Hall Effect Measurement Sequence – R±BC,AD
Current Off, B-Field Off
Short Card – Close 17, 25, 26, 27, 28, 29, 30
Prepare Measurement – Close 22
Prepare Measurement – Open 17, 25, 26, 27, 28, 30
Current On, B-Field On, Delay, Measure Voltage
Current Reverse, Delay, Measure Voltage
B-Field Reverse, Delay, Measure Voltage,
Current Reverse, Delay, Measure Voltage
Current Off, B-Field Off
Short Card – Close 17, 25, 26, 27, 28, 30
Finish – Open 22